## Slide 1: Critical Current Probe
Final
April 2012
Amy Eckerle
Andrew Whittington
Philip Witherspoon
Team 16
NHMFL
Applied Superconductivity Center

## Slide 2: Overview
Background
Introduction
  The Project
  Objectives
  Constraints
Concepts
Final Design
Project Plan
Manufacturing and Assembly
Economics
Results
Environmental Issues

## Slide 3: Super Conductivity
Superconductivity is a phenomenon that allows an indefinite amount of current to flow with negligible resistance.

## Slide 4: Critical Current
Critical Current Probes measures how much current a sample can take before it becomes non-superconducting
  Temperature
  Magnetic Field
  Orientation

## Slide 5: Samples
Sample are made of YBCO
  HTS (High Temperature Superconducting)
Samples are mounted at the base of two current leads running the length of the probe
Submerged in cryogenic liquid
Notes: Mention how this is different from LTS
Mention how the bottom of the probe is at lower temperature to allow for superconductivity which will be explained in detail later

## Slide 6: The Project
Design a critical current probe to conserve the amount of liquid helium used during a critical current measurement test
$5/liter of liquid He
Current Leads
Helium level
Cryostat
Voltage tap
Magnet
Sample
Voltage tap wire
Stainless Steel Jacket
Notes: Mention what a critical current measurement is

## Slide 7: Objectives
Conserve Helium
Test 6-8 short samples
Test one spiral sample
Deliver 1000 Amps to the samples
Durability

## Slide 8: Constraints
Constraints
  Budget of $4000
  Constant casing diameter
  Minimum length to reach center of cryostat
Current Leads
Cryostat
Stainless Steel Jacket

## Slide 9: Concepts
Heat Exchanger
HTS Leads and support
Number of Leads
Fins
Gas Insulation
Jacket Design
Ways to Reduce Helium Consumption

## Slide 10: Concept Selection
Concept Selection Table |  |  | 
Concepts | Accepted | Helium Savings per test (L) | Final
Heat Exchanger | No | --- | No
HTS Lead and Support | Yes | ≤ 26% | Modified
Number of Leads | Yes | ≤ 22% | Yes
Fins | No | 9% | Yes
Gas Insulation | No | --- | No
Jacket Design | Yes | .2% | Modified

## Slide 11: Solutions
HTS Leads and support
Low thermal conductivity with high electrical conductivity
Great reduction in copper surface area
Prevents copper leads from entering liquid helium bath
G10 Support
HTS Lead

## Slide 12: Solutions
Number of Leads
  Leads are major heat leak
  Maintain 8 samples with 6 leads
Current Leads

## Slide 13: Solutions
Jacket Design
Current Leads
Stainless steel portion
G10 portion
  interrupts thermal conduction of the stainless steel tube

## Slide 14: Final Design

## Slide 15: Final Design – Current Connects

## Slide 16: Final Design – Current Connector

## Slide 17: Final Design - HTS

## Slide 18: Final Design - Sample Holder/Guide

## Slide 19: Final Design

## Slide 20: Final Design

## Slide 21: Project Plan

## Slide 22: Manufacturing and Assembly
Machining
  ASC shop
  NHMFL shop
  Exotic Machining, Inc
Assembly
  Top flange
  Stainless steel tube
  Spacer
  Copper leads
  Current connects
  G10 connector
  Sample holder
  Soldering of HTS leads
  HTS support system
  Voltage taps

## Slide 23: Manufacturing and Assembly
Assembly
  Top flange
  Stainless steel tube
  Spacer
  Copper leads
  Current connects
  G10 connector
  Sample holder
  Soldering of HTS leads
  HTS support system
  Voltage taps

## Slide 24: Soldering of HTS Leads
Special device was created to solder 8 HTS tapes together
Soldering to copper leads using various heating methods

## Slide 25: Economics
Budget of $4000
material | quantity | cost
110 Alloy Copper rods | 6 | $291.00
G-10 Tube | 1 | $611.04
G10 Plate | 1 | $88.24
Stainless steel plate | 1 | $81.49
G10 rod | 1 | $125.77
G-10 plate | 1 | $67.40
90deg angles steel | 1 | $44.06
Sockets threaded, current connects | 6 | $180.00
copper plate | 1 | $60.95
copper bar | 1 | $62.52
G-10 Plate | 1 | $60.18
aluminum bar | 1 | $49.34
cartridge heater | 7 | $247.24
cartridge heater | 2 | $40.52
wing nut | 1 | $11.21
compression springs | 2 | $26.72
Exotic Machining | 1 | $400.00
 | Total | $2,447.68

## Slide 26: Preparing
Applied voltage taps along the probe
Soldered a strip of HTS to complete circuit
Covered base with G-10 to protect HTS

## Slide 27: Testing
Took place the NHML
Lifted by crane into cryogenic bath
Preliminary testing
  To make sure HTS Lead was superconducting
Final testing in Liquid helium to measure burn off
Notes: Mention preliminary testing was don in liquid nitrogen
Current Leads 4 and 5

## Slide 28: Environmental/Safety Hazards
G-10
Flux
Cutting Fluid
Machining Hazards
Heating Hazards
Handling Liquid Helium
Electric Current
Weight
G-10
Flux

## Slide 29: Results
Preliminary test results (in liquid nitrogen)
  Took place after 12pm 4/5/12
Future testing (in liquid helium)
  Will take place at NHMFL 4/6/12

## Slide 30: Conclusions
Testing in liquid nitrogen will completed on two leads
Full testing will be completed on 4/5/12 if the preliminary is successful

## Slide 31: Acknowledgments
Dr. Hovsapian, Adjunct Faculty, Florida State University, Mechanical Engineering, Ph.D.
Dr. Kosaraju, Adjunct Faculty and Postdoctoral Researcher
Dr. Hellstrom, Ph.D. Materials Science, Stanford University,
Dr. Trociewitz, Associate Scholar/Scientist, ASC
Applied Superconductivity Center
NHMFL
Bill Sheppard, NHMFL Machine Shop
Robert Stanton, NHMFL
Bill Starch, ASC Machine Shop
Dimitri Argonaut, ASC Machine shop

## Slide 32: References
Çengel, Yunus A., Robert H. Turner, and John M. Cimbala. Fundamentals of Thermal-fluid Sciences. Boston: McGraw-Hill, 2008. Print.
Ekin, Jack W. . Experimental Techniques for Low-temperature Measurements. New York: Oxford UP, 2006. Print.
Thomas, Lindon C. Fundamentals of Heat Transfer. Englewood Cliff, NJ: Prentice-Hall, 1980. Print.

## Slide 33: Drawings