## Slide 1: Tabletop Torsion Device
Final Presentation
By: Brendan Keane, Logan McCall, Reggie Scott, Mark Swain
Instructor: Dr. Nikhil Gupta
Sponsor: Dr. Philip Flater (Air Force Research Laboratory Eglin AFB)
Advisor: Dr. Simone Hruda

## Slide 2: Overview
Introduction
Design and Analysis
Prototype Testing
Project Management
Conclusion
Reginald Scott

## Slide 3: Introduction
Current torsion tester at the Air Force Research Laboratory (AFRL) Munitions Directorate at Eglin AFB is approximately 5 meters in length
Small sample geometries result in inaccuracy and inefficiency
Figure 2.1: Existing Torsion Machine
“DISTRIBUTION A: Approved for public release, distribution unlimited. (96ABW-2014-1649)”
Reginald Scott

## Slide 4: Specimen Geometry
Small specimens are a result of original plate stock dimensions
Specimens tested:
  Aluminum 2024
  Titanium 64
Original plate stock
Blank removed
Sample machined
Figure 3.1: Specimen production
Dimension | Measurement (mm)
Total Length | 58.4
Gauge Length | 12.7
Width | 14.3
Inner Diameter | 9.09
Fillet Radius | 27.9
Grip Length | 20
Table 3.1: Specimen dimensions
“DISTRIBUTION A: Approved for public release, distribution unlimited. (96ABW-2014-1649)”
Reginald Scott

## Slide 5: Design Requirements
Maximum torque output: 100Nm
Machine must fit in an area (footprint) of 1m2
Monotonic (one direction) free-end torsion loading
Free end must allow for axial motion
Compatible with measurement equipment used by AFRL
Budget - $2,000
Reginald Scott

## Slide 6: Design Requirements
Goal Statement:
Design a more effective way of testing small specimens in free-end torsion
Reginald Scott

## Slide 7: DIC (Digital Image Correlation)
“DISTRIBUTION A: Approved for public release, distribution unlimited. (96ABW-2014-1649)”
Figure 6.2 Example DIC use in torsion sample
Figure 6.1 DIC setup at Eglin AFRL
Brendan Keane

## Slide 8: Load Measurement
Current method used by AFRL to measure applied stress:
  Strain gage placed on elastic bar
  Use shear modulus relationship to output stress
Figure 7.1: Stress vs. strain plot showing the shear modulus relationship
G =
Brendan Keane

## Slide 9: Breakdown of Design
Torsion Tester
Load Generation
Manual Power
Load Measurement
Linear Motion
Compatibility
Gripping Mechanism
Load Application
Friction Reduction
Motor & Transmission
Frame
Material Selection
User Safety
Brendan Keane

## Slide 10: Load Generation
Design | Cost | Weight | Accuracy | Complexity | Maintenance | Variability | Total
Weight Factor | 0.25 | 0.05 | 0.25 | 0.1 | 0.1 | 0.25 | 
(1) Crank System | 5 | 3 | 1 | 5 | 5 | 1 | 2.9
(2) Hydraulic | 1 | 1 | 5 | 1 | 1 | 5 | 3
(3) AC Motor | 3 | 3 | 5 | 3 | 3 | 5 | 4
(1)
(2)
(3)
Brendan Keane

## Slide 11: Load Application
Design | Cost | Weight | Reliability | Complexity | Variability | Total
Weight Factor | 0.25 | 0.15 | 0.3 | 0.1 | 0.2 | 
(1) 3-Jaw Chuck | 5 | 5 | 5 | 3 | 5 | 4.8
(2) 4-Jaw Chuck | 3 | 5 | 5 | 3 | 1 | 3.5
(3) Self-Aligning Vise | 3 | 1 | 3 | 5 | 5 | 3.3
(4) Collet | 5 | 5 | 5 | 3 | 1 | 4
(1)
(2)
(3)
(4)
Brendan Keane

## Slide 12: Linear Motion
Design | Cost | Weight | Durability | Complexity | Total
Weight Factor | 0.4 | 0.2 | 0.2 | 0.2 | 
(1) 4 Rail Ball Bearing Guide | 1 | 2 | 1 | 1 | 1.4
(2) 2 Track Roller Bearing Guide | 3 | 5 | 5 | 5 | 4.2
(3) 2 Rail Ball Bearing Guide | 5 | 5 | 3 | 3 | 4.2
(1)
(2)
(3)
Brendan Keane

## Slide 13: Frame
Design Considerations:
  Strength
  Machinability
  Weight
  Cost
Material Selection:
  Steel
Figure 12.1: Frame design
Brendan Keane

## Slide 14: Concept Evolution
Brendan Keane

## Slide 15
Figure 5.1: CAD representation of design
Design Overview
Logan McCall
Steel Shaft
32”
13”

## Slide 16: Load Generation: AC Gearmotor
Vendor: Grainger
Rpm: 18
Gear ratio: 95:1
Max torque: 116 Nm
Weight: 26 lbs
Figure 15.1: AC Gearmotor
Logan McCall

## Slide 17: Motor Control: VFD
Vendor: Automation Direct
Single phase input, three phase output
Digital keypad
Forward and reverse
Expandable
Variable Frequency Drive (VFD)
Figure 16.1: VFD keypad
Logan McCall

## Slide 18: Load Application: 6-Jaw Chuck
Vendor: LittleMachineShop
Weight: 6.4 lbs
Outer diameter: 3.94 in (10 cm)
Inner diameter: 1.02 in (2.59 cm)
Figure 17.1: 6 Jaw Chucks holding specimen
Logan McCall

## Slide 19: Load Measurement
Coupler material: Al 6061
Coupler properties
  Shear modulus: 26 GPa
  Shear strength: 76 MPa
  Outer diameter: 1.75 in (4.45 cm)
  Inner diameter: 0.75 in (1.90 cm)
Figure 18.1: FEA performed on coupler showing strain developed under a simulated load
Logan McCall

## Slide 20: Linear Motion: 2 Rail Ball Bearing Guide
Vendor: Grainger
½ inch steel rails
Slot design to ensure alignment
Figure 19.1: Linear guide system
Logan McCall

## Slide 21: Frame Design
Steel frame
Shape: hollow square shaped cross-section
Dimensions:
  1/8 in inner supports – 304 stainless steel
  1/8 in outer frame –
  low carbon 1015 steel
Figure 20.1: CAD drawing of frame (bottom view)
Logan McCall

## Slide 22: Final Design
Logan McCall

## Slide 23: Prototype Testing
Mark Swain

## Slide 24: Testing Results
Tester was successful in breaking aluminum specimen
Fracture was at center
During initial testing, free end chuck had slight rotation
  Solution: Force fit new connecting rod into free end
Cylindrical grips slipped during initial testing
  Solution: Use hex-grip specimen or increase friction on cylindrical specimen
Figure 23.1: Fractured specimens
Mark Swain

## Slide 25: Budget
Chart: 
| Load Measurement | Load Generation | Load Application | Linear Motion | Housing | Miscellaneous | Travel
 | 0 | 816.55 | 458.94 | 331.48 | 181.9 | 50 | 0
Budget: $2,000
Spent: $1,844
Net: + $156
Mark Swain

## Slide 26: Project Progression
Fall
Background research
Design breakdown
Component research
Design analysis
Optimal part selection
Budget breakdown
Purchase orders
Spring
Delivery of parts
Machining
Assembly
Motor/VFD testing
Prototype testing
Troubleshooting
Final assembly
Mark Swain

## Slide 27: Conclusion
Maximum torque output: 100Nm
Machine must fit in an area (footprint) of 1m2
Monotonic (one direction)free-end torsion loading
Free end must allow for axial motion
Compatible with measurement equipment used by AFRL
Budget - $2,000
Mark Swain

## Slide 28: Conclusion
The team was successfully designed and manufactured a tabletop torsion tester for AFRL
What we learned:
  Focus on the main customer requirements
  Consult as many resources as possible
  Teamwork is essential
  Unforeseen circumstances
Figure 27.1: Final assembly
Mark Swain

## Slide 29: Questions?

## Slide 30: References
Carter, B. (2008). Texas Instruments: Op Amp Noise Theory and Applications. Retrieved September 22, 2014
Flater, P. (2014). Tabletop Torsion Test. Eglin, FL: Air Force Research Laboratory.
Ilic, M. (2014, October 11). Clamp for centering. Retrieved from GRABCAD: https://grabcad.com/library/stega-za-centriranje-clamp-for-centering-1
Lathe Chuck. (2014, October 7). Retrieved from GRABCAD: https://grabcad.com/library/lathe-chuck-3
Linear Motion Systems. (2014, September 28). Retrieved from Stock Drive Products: https://sdp-si.com/eStore/coverpg/linearmotion.htm
Mark Swain

## Slide 33: Fall Gantt Chart

## Slide 34: Spring Gantt Chart

## Slide 35: Specimen Dimensions
Brendan Keane
Units in inches

## Slide 36: Appendix