## Slide 1: Catheter Rotation Measuring Device
Madison Burke | Savannah Buro | Amari Deleon | Christopher Hayot
Alberto Villacrez | Jake Anthony Vickers
Notes: Customize your title slide to be appealing to the audience. This is the first thing they see apply some graphic design elements that may summarize your presentation.

## Slide 2: Team Members
Departments of Mechanical & Biomedical Engineering
Systems & Testing Engineer
(ME)
Bioinstrumentation Engineer
(BME)
Software & Design Engineer
(BME)
Biomaterials & Quality Engineer
(BME)
Biomaterials & Testing Engineer
(BME)
Software & Quality Engineer
(BME)
Jake Vickers
Alberto Villacrez
Christopher Hayot
Amari Deleon
Savannah Buro
Madison Burke

## Slide 3: Sponsors & Advisors
Dr. Stephen Arce
FAMU-FSU
Dept. of CBE
Amar Patel
Biosense Webster
Chuck Lindholm
Biosense Webster
Departments of Mechanical & Biomedical Engineering
Dharani
Balakumar
Biosense Webster
Notes: SOMEBODY WHO SPEND MONEY AND TIME

## Slide 4: Project Description
To design and construct a measuring device that confirms the rotational relationship between a cardiac catheter’s handle and tip, allowing design engineers and physicians increased control precision.
Departments of Mechanical & Biomedical Engineering
Tip
(Distal End)
Handle
(Proximal End)

## Slide 5: Cardiac Catheter Background
Catheter Insertion Sites
Departments of Mechanical & Biomedical Engineering
Notes: Uses for Cardiac Catheter and why our device is need.
Where does is go, what does it do, how does it work, why we need . Copper wire anaolgy

## Slide 6: Progress
Departments of Mechanical & Biomedical Engineering

## Slide 7: Initial Design
Departments of Mechanical & Biomedical Engineering
Arduino
LCD Display
MASACath
Handle Angle: 65.1 deg
Tip Angle: 43.2 deg
Arduino
Magnetic Encoder Module
IMU Module
Magnet
Magnet Tube
Guiding Tube
AS5600
Baseboard
Rigid Track

## Slide 8
UI Display
Arduino
Handle
Clamps
IMU
Sheath
Guiding Tube
Tip
Motor
Peg Holes
Rigid Supports
Departments of Mechanical & Biomedical Engineering
Wood Baseboard
Plexiglass Cover

## Slide 9: Handle Module
Departments of Mechanical & Biomedical Engineering
Clamps
Connects motor shaft, IMU, and handle as rigid body
Two points of contact with handle
M3 screws & washers
Slots hold stepper motor attachment
IMU mounted on top
Handle
IMU
Stepper Motor
IMU (Internal Measurement Unit)
Measures input rotation from stepper motor
Motor
NEMA 17 Stepper Motor
Reads & outputs rotation, defined by operator

## Slide 10: Handle Module
Departments of Mechanical & Biomedical Engineering
IMU
Motor
Clamps
Purpose: To provide a set input rotation to use as a reference for measured tip rotation.

## Slide 11: Tip Module
Departments of Mechanical & Biomedical Engineering
Magnet
Magnet Tube
Guiding Tube
Magnetic
Encoder
Magnetic Encoder
Measures rotation translated from the handle
Locking mechanism to prevent catheter from slipping out
Tight Fit
Locking
Mechanism
Magnet
Guiding Tube
Magnet Tube

## Slide 12: Environment Module
Departments of Mechanical & Biomedical Engineering
Test Enclosure
Rigid box with sliding plexiglass lid
Enclosure for environmental control
Baseboard has peg holes for easy setup between different orientations
Current Enclosure Build

## Slide 13: Environment Module
Departments of Mechanical & Biomedical Engineering
Component Interfaces
Catheter guided through the sheath, with tip free to move at end.
Sheath is guided through rigid frame attached to baseboard.
Sheath
“Vein”
Cross-Section of V mount
Catheter
Peg Hole
Catheter
Sheath
“Vein”
Cross-Section of Tube Shaft
Peg Hole
Side View of Interface Connection

## Slide 14: Benchmark Testing Schedule
Departments of Mechanical & Biomedical Engineering
Pencil Test
Purpose:
Verify handle & tip sensors are properly calibrated
Method:
Treat catheter as a rigid body
Perform 5+ measurements
Expected Result:
1:1 rotation translation between two ends

## Slide 15: Benchmark Testing Schedule
Departments of Mechanical & Biomedical Engineering
Gage R&R Study
Purpose:
Validate repeatability and reproducibility
Test setup constraints & measurement system are accurate and precise
Method: “Crossed GR&R Study”
3 operators, 1 catheter
5+ orientation measurements
Expected Results:
<10% variation of measured angle between operators & orientations
Computer Vision Validation
Purpose:
Validate tip sensor design choice
Method:
Perform “Pencil Test” but with computer vision MATLAB analysis at tip end
Using “PENTAWAY” catheter, mark each tip with a different color for observation
Expected Results:
Accuracy of: MATLAB angle reading ≤ magnetic encoder

## Slide 16: Current Work & Issues
Departments of Mechanical & Biomedical Engineering
Environment Simulation
Testing Design
CAD Design
Enclosure Temperature control
Vein/sheath pressure simulation
Addition of fluid between sheath and vein structure
Handle & motor clamps
Tip & UI module housings
“Vein” structure
Enclosure & Pegboard Modeling
Construction of Pencil Test setup
Benchmarks of successful test
Computer vision programming

## Slide 17: Project Description
To design and construct a measuring device to confirm the rotational relationship between cardiac catheters’ handle and tip ends, allowing physicians and design engineers to increase control precision.
Departments of Mechanical & Biomedical Engineering
Tip
Handle
Electrodes

## Slide 18: Background
Customer Needs
Project Scope
Aug-Sep
Concept Generation & Selection
Functional Decomp & Targets
Oct
510(k) Proposal
Prototyping & Validation Testing
Nov-Dec
Departments of Mechanical & Biomedical Engineering

## Slide 19: Progress
Departments of Mechanical & Biomedical Engineering

## Slide 20: Customer Needs
Departments of Mechanical & Biomedical Engineering
1
2
3
4
Need Statements
Surgeons Designers
Need Interpretation
Reflection on
Results & Process
Raw Data Gathering
Angle Accuracy Analysis
Notes: What is the current problem for the catheter?
What do you expect the device to do?
Dynamic process of identifying customer need
Raw Data Gathering
Weekly lectures
Cardiac ablation lab
Surgery Shadowing
Need statement
Doctors: It bothers me that when I rotate the catheter handle, the tip doesn’t rotate the same amount. I want to know the angle of rotations.
Designers: We want to have a device to measure the rotational angle of the catheter tip/handle so that we could determine why the current translation is not direct and use it in the development process of catheters.
Controls are precise
Results are reliable
Easy to install and use
Interpret needs
The device measures rotation angles of the catheter tip/handle.
The device can be used for testing purpose.
The device can be used to explore relationships between catheter tip and handle rotations.
The device can produce reliable results and can be precisely controlled.
TO know the rotation angle, to have accuracy, and to analyze the relationship

## Slide 21: Key Goals
Departments of Mechanical & Biomedical Engineering
Accurate
± 1 degree
minimum accuracy
Repeatable
Withstands
multiple uses
Quantifiable measuring
system
Translatable across different types of catheters and operators

## Slide 22: Assumptions
Device is suitable for only one catheter at a time.
Departments of Mechanical & Biomedical Engineering
Catheters are assumed to be manufactured by Biosense Webster.
Device not used inside human body
1
Biosense
Webster

## Slide 23: Catheter Motion
Departments of Mechanical & Biomedical Engineering
Body Rotation
Linear Translation
Tip Deflection
Tip Deflection
Body Rotation
Linear Translation
Notes: Before we introduce functional decomp, let’s take a look at how catheter works

## Slide 24: Functional Decomposition
Departments of Mechanical & Biomedical Engineering
Ease of Use
User
System
Measurement
Sensors
Controls
21.CFR.820

## Slide 25: Critical Targets
Departments of Mechanical & Biomedical Engineering
Ease of Use
Measurement
System
User
Simple Test Procedure
            < 10 minutes to completion
    Automatic Data Processing
      ~2 minutes
      Interactable User Interface
        Digital Display
        Formatted Results
Easily Accessible
  Length doesn’t exceed longest catheter by > 5 inches
  Mounted on top of common surface
Compatible with varying Catheter Models
< 5 setup steps needed to change between models

## Slide 26: Critical Targets
Departments of Mechanical & Biomedical Engineering
Ease of Use
Measurement
Sensors
Measure Tip Deflection
resolution w/in ± 1mm
Measure Shaft Rotation & Bend
resolution w/in ± 1 degree
Controls
Separates Proximal & Distal Ends
2 Sensor Systems
Compare Proximal & Distal End Data
Microcontroller
Error Analysis
  Noise Controller
21.CFR.820
Simulates In Vitro Environment
Sheath Diameter < 1 mm + Vein Diameter
Quality Schedule
Re-Calibration: > 6 months
Sterilization: after every use

## Slide 27: Concept Generation
Departments of Mechanical & Biomedical Engineering
Physical Quantity Input
Catheter Measuring Device
Angle
Torque
Signal Output
External Search
Internal Search
Engineer Experts
Catheter Patents
Benchmarking
Individual Brainstorming
Team collaboration

## Slide 28: High Fidelity
Departments of Mechanical & Biomedical Engineering

## Slide 29: Rotary Angle Sensor
Departments of Mechanical & Biomedical Engineering
Optical Encoder
Proximity Sensor
Computer Vision

## Slide 30
Departments of Mechanical & Biomedical Engineering
Medium Fidelity

## Slide 31: CATHerine
Departments of Mechanical & Biomedical Engineering
Arduino
Rigid Track
Torque
Meter
Servo
Motor
Custom Connector
Polyurethane
Sheath

## Slide 32: Sauron
Departments of Mechanical & Biomedical Engineering
Arduino
LCD Display
MASACath
Handle Angle: 65.1 deg
Tip Angle: 43.2 deg
Arduino
Magnetic Encoder Module
IMU Module

## Slide 33: Handle Measurement
Departments of Mechanical & Biomedical Engineering

## Slide 34: Tip Measurement
Magnet
Magnet Tube
Guiding Tube
AS5600
Departments of Mechanical & Biomedical Engineering

## Slide 35
Magnet
Magnet Tube
Guiding Tube
AS5600
Strain Gauge
Measures dynamic torque
Measures angle
Signal Amplifier
Departments of Mechanical & Biomedical Engineering

## Slide 36: Future Work
Departments of Mechanical & Biomedical Engineering
Final Design
Material
Purchasing
Prototype
Evaluation &
Environment
Building
Verification & Validation Testing
Detailed Design
&
Data Analysis

## Slide 37: Future Work
Departments of Mechanical & Biomedical Engineering
Final Design
Material
Purchasing
Prototype
Evaluation &
Environment
Building
Verification & Validation Testing
Detailed Design
&
Data Analysis

## Slide 38: References
Mayo Foundation for Medical Education and Research. (n.d.). Catheter insertion points for cardiac ablation. Mayo Clinic. Retrieved January 25, 2023, from https://www.mayoclinic.org/tests-procedures/atrial-fibrillation-ablation/multimedia/img-20302581
Mayo Foundation for Medical Education and Research. (2022, April 30). Atrial fibrillation ablation. Mayo Clinic. Retrieved January 25, 2023, from https://www.mayoclinic.org/tests-procedures/atrial-fibrillation-ablation/about/pac-20384969#:~:text=During%20cardiac%20ablation%2C%20catheters%20are,the%20groin%2C%20shoulder%20or%20neck
Departments of Mechanical & Biomedical Engineering

## Slide 39: Backup Slides
Departments of Mechanical & Biomedical Engineering

## Slide 40: Primary Market
& Biomedical Engineering

## Slide 41: Secondary Markets
Medical
Schools
Research Labs
&
Institutions
Catheter Manufacturers

## Slide 42
Departments of Mechanical & Biomedical Engineering
Arduino
LCD Display
MASACath
Handle Angle: 52.2o
Tip Angle: 38.4o
Arduino
Magnetic Encoder Module
IMU Module
Polyurethane
Sheath

## Slide 43: Concept Classification Tree

## Slide 44: House of Quality
Departments of Mechanical & Biomedical Engineering

## Slide 45: Concept Selection
Departments of Mechanical & Biomedical Engineering
Learning
Rapid Prototyping
House of Quality
Pugh Chart
Selected Concept
Prototyping
Testing & Evaluation
House of Quality
Pugh Chart
Customer Need Development
Selected Concept

## Slide 46: Surgical Environment
Notes: Show if ppl ask during question section

## Slide 47
Departments of Mechanical & Biomedical Engineering
Sauron Software System
Sub System 1
Microcontroller
Module
LCD
Module
IMU
Module
Magnetic
Encoder
Module
Item 1
Item 2
Item 3

## Slide 48: Functional Decomp
Repeatability
Sensors
Varying Test Setups
Data Processing
Distinguishes Proximal/Distal Ends
Measure Angles of Bending & Twist
Measure Tip Displacement
Compatible with different catheter models
Compact Design
Automatic Data Processing
Identifies locations of translation drop-offs
Accounts for varying positional set-ups

## Slide 49: Functional Decomposition
Departments of Mechanical & Biomedical Engineering
Ease of Use
Digital Display
Compact Design
User
Adjustable System Parameters
Repeatable
System
Measurement
Tip Rotation & Displacement
Rotation of Handle
Sensors
Distinct Proximal/Distal End Containers
Signal Transducing & Processing
Controls
Regulatory
Quality & Calibration Schedule
Sterilized Environment
21.CFR.820

## Slide 50: IMU for Catheter Handle
9-axis IMU ( 3-axis gyro + acc + mag) + Arduino Mega 2560
Noise Filter (Kalman VS Complementary Filter) Programming
Vpython 3D Modeling, real-time visualization
Why IMU?
Freedom for handle movement
Future modification (Remote Catheter System,
Catheter Robotics, Gesture machine learning)
& Biomedical Engineering