## Slide 1
Stemless Reverse Shoulder Arthroplasty
Valeria Aguilera, Ethan Corey, William Crittenden, Alex Nowzamani, Matthew Mohammed, Nickolas Wolniewicz
Design a reverse stemless shoulder implant to improve range of motion, strength in the shoulder and reduce overall pain. Create testing procedures to ensure the implant would not fail.
Background
About 50,000 people undergo shoulder replacement surgery annually.
There are over 800,000 people in the U.S. with a shoulder implant.
Reverse Stemless implants may reduce bone loss, protect rotator cuff, and lower fracture risk.
Testing variables like fin count can help develop future implant designs.
Mechanical Testing
Shear
Lever Out
COMSOL Model
Bone Substitute
Bone Substitute
Rotator Cuff
Deltoid
Testing was split into two sections:
(1) Measuring force with a 10 kg Force Meter​ on Extruded Polystyrene Insulation
(2) Measuring force with a Digital Force Gauge and 10 and 15 PCF Bone Blocks.
Chart: Maximum Lever Out Forces
| 2 Horizontal | 3 fins Low | 4 fins | 5 fins | 6 fins
2 Horizontal | 411.1
3 | 291.975
3 Low | 245.3
4 fins | 228.0333
5 fins | 295.7333
6 fins | 375.6
*Low = fins offset 30° from original
High = fins offset starting at 12 o clock position
Chart: Maximum Shear Out Forces
| 2 Horizontal | 3 fins Low | 4 fins | 5 fins | 6 fins
3 Low | 432.8
3 Original | 500.55
3 High | 637.1
*Low = fins offset 30° from original
High = fins offset starting at 12 o clock position
(2)
Discussion
Acknowledgements
Advisor: Tom Vanasse​
Instructors: Stephen Arce, Ph.D. Shayne McConomy, Ph.D.​
Equipment Use: FAMU-FSU College of Engineering, Formlabs, Emily Hubiki
Chart: Maximum Lever Out Force
2 | 55.4756
3 | 58.5068
4 | 56.0641
5 | 54.3965
6 | 44.7826
(1)
Applying a compressive force to walls of the bone block increased the peak force values the implants could withstand before failure.
The 15 PCF bone blocks were more susceptible to fractures forming during the implantation process as opposed to the softer 10 PCF bone blocks.
When the fins were orientated perpendicular to the applied force, the peak force performance was higher
When a fin was oriented in the rotational path of the cage, the force was lower
Limited by the number of bone blocks
Prototyping
0 N
100 N
200 N
400 N
500 N
300 N
Shear Stress (N/m2)
Force Applied
Future Plans
Improving COMSOL models
Improving tests to make them more robust by adding metal plates to the bone block walls
Realistic prototyping
  Materials: Titanium and Cobalt Chrome alloy, HDPE
  Machinery: SLS
  Testing Mechanisms: Instron
Results
Objective
200 N
0 N
100 N
500 N
400 N