OMEGASCOPE THERMOMETER: Tests and Calibration Experiment #2

- Calibration

The objective of this section was to calibrate the IR thermometer by determining the temperature of known surfaces.  The temperature of the control surface was determined by using a RTD.  Four different surfaces were used: black, white1, white 2 and polished aluminum.

Experimental Procedure:

- Use the Thermal Radiation Cube, Leslie's Cube and set the meter to the desired temperature setting.

- Connect a multimeter to the cube.  Set the multimeter to calculate electrical resistance.

- Read the resistance of the meter.  Compare this value with the table on the user's manual.  Determine the correct surface temperature.  .

- Adjust the emissivity on the IR thermometer.

- Determine the surface's temperature using the IR thermometer.

- Calculate the difference between both readings and draw a calibration curve.

Behind: White 1

White 2

Polished Aluminum

Black

1st Calibration: Determining the Unknown Emissivites

The objective of this experiment was to determine the emissivities of each of the four different surfaces.  The cube was set to a predetermined temperature, and then the emessivity on the scope was adjusted until the temperature reading matched the actual temperature.  Table 1.1 contains the experimental results.

|  |  |  |  | Surface Emissivity |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Temp. Set | Resistance (K Ohms) | Temp. (deg. C) | B | P | W1 | W2 |
| 1 | 106.6 | 23.3 | 1 | 1 | 0.94 | 0.47 |
| 2 | 105.6 | 23 | *1 | *1 | 1 | 0.96 |
| 3 | 85.7 | 29 | 0.91 | 1 | 0.85 | *1 |
| 4 | 50.3 | 39 | 0.78 | 0.22 | 0.96 | 0.15 |
| 5 | 23.25 | 59 | 0.94 | ** | 0.96 | 0.2 |
| 6 | 11.63 | 79 | 0.93 | ** | 0.95 | 0.2 |
| 7 | 6.86 | 94 | 0.92 | ** | 0.95 | 0.2 |
| 8 | 4.21 | 109 | 0.9 | ** | 0.95 | 0.2 |
|  |  | Avrg. Emissivity | 0.896667 | ** | 0.945 | 0.19 |

Table 1.1 lists the experimental results for test #1

After performing the first test, the emissivity for the side painted black was 0.89.  This value differed from the one given on table B-2 of the user's manual.  For the second surface, White 1 (W1), the average emissivity was 0.945.  For W2 the emissivity equaled 0.19.  W2 proved to be the most constant temperature.  As it can be seeing from the table, it was impossible to determine its emissivity.  From table B-2, the emissivity for this material was listed as 0.09.  The minimum emissivity setting in the scope was 0.1.  As a result, these temperatures were impossible to measure with a high level of accuracy.

2nd Calibration: Determining the Surface Temp. With Known Emissivites

The purpose of this section was to determine the temperature of different surface finishes.  This test was performed to two surfaces: black and W1.  The procedure was the following:

- Set a desired temp.

- Set the emissivity on the scope to either 0.96 for black or 0.95 for W1.

- Determine surface temp.

- Change the emissivity in order to match the real surface temp.

Table 1.2 lists the experimental results for the second calibration experiment.

| Temp. Set | Resistance (K Ohms) | Temp. (deg. C) | B=0.96 | Real e | W1=0.95 | Real e |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 88.8 | 27 | 27 | 0.96 | 27 | 0.95 |
| 2 | 72.5 | 32 | 31 | 0.93 | 31 | 0.96 |
| 3 | 55.1 | 38 | 37 | 0.93 | 38 | 0.95 |
| 4 | 34.8 | 49 | 47 | 0.94 | 49 | 0.95 |
| 5 | 22.49 | 60 | 59 | 0.95 | 60 | 0.95 |
| 6 | 6.74 | 94 | 92 | 0.92 | 94 | 0.95 |
| 7 | 4.08 | 105 | 105 | 0.96 | 108 | 0.93 |
| 8 | 3.22 | 118 | 113 | 0.9 | 117 | 0.93 |
|  |  | Avg. Emissivity |  | 0.93625 |  | 0.942 |

Table 2.1

After all the data points were tabulated figures 2.1 and 2.2 were drawn.  These figures show how closely the temperature readings for both surfaces resembled that of the actual temperature.

From the figure above it could be seeing that the temperature calculated by the scope using the two different emissivities was very close to the actual temperature.  As a result, all three-temperature lines lay on top of each other.  This example shows the accuracy of the IR thermometer.  The actual calibration line is displayed in figure 1.2.

Figure 2.2