OMEGASCOPE THERMOMETER: Determining the Relationship between Emissivity and Temp .

Experiment #3

The objective of this test was to prove that the temperature varies linearly with the emissivity.  During experiment #1, it was proven that the temperature and emissivity had an inverse relation.  This meant, that as the emissivity increased, the temperature calculated by the scope decreased.  Now, we want to know the nature of this relationship.

- Set the temp. of the cube to a fix value.

- Set the emissivity on the scope to its minimum value, 0.1.

- Record the temperature reading.

- Increase the emissivity by 0.05.

- Take a second temperature reading.

- Repeat this until the emissivity equals 1.

Table 1.1 and 1.2 lists the experimental results for the third calibration test.

| e | 1/e | Temp. | Temp K |
| --- | --- | --- | --- |
| 0.1 | 10 | 467 | 740 |
| 0.15 | 6.666667 | 355 | 628 |
| 0.2 | 5 | 288 | 561 |
| 0.25 | 4 | 252 | 525 |
| 0.3 | 3.333333 | 225 | 498 |
| 0.35 | 2.857143 | 206 | 479 |
| 0.4 | 2.5 | 189 | 462 |
| 0.45 | 2.222222 | 176 | 449 |
| 0.5 | 2 | 166 | 439 |
| 0.55 | 1.818182 | 156 | 429 |
| 0.6 | 1.666667 | 149 | 422 |
| 0.65 | 1.538462 | 141 | 414 |
| 0.7 | 1.428571 | 134 | 407 |
| 0.75 | 1.333333 | 129 | 402 |
| 0.8 | 1.25 | 124 | 397 |
| 0.85 | 1.176471 | 119 | 392 |
| 0.9 | 1.111111 | 115 | 388 |
| 0.95 | 1.052632 | 112 | 385 |
| 1 | 1 | 108 | 381 |

Table 1.1

| e | 1/e | Temp. | Temp K |
| --- | --- | --- | --- |
| 0.1 | 10 | 479 | 752 |
| 0.15 | 6.666667 | 363 | 636 |
| 0.2 | 5 | 298 | 571 |
| 0.25 | 4 | 259 | 532 |
| 0.3 | 3.333333 | 232 | 505 |
| 0.35 | 2.857143 | 211 | 484 |
| 0.4 | 2.5 | 194 | 467 |
| 0.45 | 2.222222 | 180 | 453 |
| 0.5 | 2 | 169 | 442 |
| 0.55 | 1.818182 | 159 | 432 |
| 0.6 | 1.666667 | 151 | 424 |
| 0.65 | 1.538462 | 144 | 417 |
| 0.7 | 1.428571 | 137 | 410 |
| 0.75 | 1.333333 | 132 | 405 |
| 0.8 | 1.25 | 127 | 400 |
| 0.85 | 1.176471 | 122 | 395 |
| 0.9 | 1.111111 | 118 | 391 |
| 0.95 | 1.052632 | 114 | 387 |
| 1 | 1 | 110 | 383 |

Table 1.2

As it can be seeing tables 1.1 and 1.2 lists the data point for this experiment.  As it was mentioned above, the objective of this experiment was to identify the nature of the relationship between the emissivity and temperature.  From the Stefan-Boltsmann Law, we can determine this connection theoretically.   If we take this equation and solve for q/e we obtain the following expression:

qε=σ(T4−Tatm4)

From this equation it is evident that the relationship between 1/ and T is linear.  This idea is bolstered by the experimental data.  Figures 1.1 and 1.2 illustartes this behavior.

Figure 1.1

Figure 1.2